3GPP TSG-RAN5 Meeting #101	R5-237138r2

# Chicago, US, 13th Nov 2023 – 17th Nov 2023

| CR-Form-v12.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| TEXT PROPOSAL |  |  |  |  |  |  |  |  |
|  | 38.561 | CR | N/A | rev | - | Current version: | 1.0.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | pCR with TRP/TRS MU Updates |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Keysight Technologies UK Ltd |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | NR_FR1_TRP_TRS-UEConTest |  |  |  |  |  | Date: |  |  | 2023-10-30 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-17 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) |  |
| Reason for change: |  | Update to measurement grids and MU |  |  |  |  |  |  |  |  |
| Summary of change: |  | MU and grids updated based on empirical results |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | TRP/TRS testing is not performed properly with the correct MUs |  |  |  |  |  |  |  |  |
| Clauses affected: |  | A.3.3.3, A.4.2.12, A.4.3.1, A.4.4.1 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Discussion Paper in R5-237137 |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | R5-237138r1: confirm the Uncertainty of network analyzer MU; compromise value for Communication Tester: uncertainty of the absolute output level MU (from R5-237152). Compromise on measurement grids based on empirical data submitted to meeting; MU determined based on legacy grids. Aligned TRS test procedure in A.3.3.3 R5-237138r2: corrected heading row in Table A.4.2.12-1; revised Combined Standard Uncertainty in Table A.4.3.1-2 |  |  |  |  |  |  |  |  |

<<< START OF CHANGES >>>

#### A.3.3.3	TRS Test procedure

The TRS of the DUT is measured by sampling effective isotropic sensitivity (EIS) of the DUT with three-dimensional scan at various locations surrounding the device. The measurement is performed with a constant sampling step in both theta () and phi () axes using any of the measurement grids and quadratures options outlined and applicable to TRS in Table A.4.2.12-1.

EIS, or receiver sensitivity measurements, is defined as the minimum downlink signal power received at the UE antenna input required to provide a data throughput rate greater than or equal to 95% of the maximum throughput of the reference measurement channel (RMC) (the maximum throughput is per Appendix A of TS 38.521-1 [TBD]).

For TRS measurement, the evaluations shall be performed at maximum transmit power.

The measurement procedure includes the following steps:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the measurement antenna.

3)	Follow steps 1 through 4 in section 7.3.2.4.2 of TS 38.521-1 [TBD], with the following exception: determine each EIS, i.e., by adjusting the downlink signal level until the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point. The downlink power step size shall be no more than 0.5 dB when the RF power level is near the NR sensitivity level.

The TRS value is calculated using the equation outlined in Clause 5.2 in [TBD]. The reported TRS value shall be corrected by -0.5·output level step resolution of the final power step search and note the correction in the test report.

This TRS test procedure is common to both SA and EN-DC measurements. The detailed UE configurations for TRS test in SA and EN-DC mode are specified in Clause A.2.

<<< Skip Unchanged Sections >>>

#### A.4.2.12	Coarse sampling grid

Editor’s Note: Further analyses of measurement grids and impact on measurement uncertainty based on empirical data are pending.

This contributor describes the uncertainty of the measured TRP/TRS value due to the finite number of measurement grid points. Decreasing of sampling density to finite number of samples affects the measurement uncertainty. Different TRP quadratures also have an effect on the MU. The grid options for TRP/TRS with associated MUs for constant-step size grids are summarized in Table A.4.2.12-1.

Table A.4.2.12-1: Measurement grid options for TRP/TRS

| Test Metric | Frequency Range | Quadrature | [°] | [°] | Min. Number of Grid Points (Note 1) | Std. Uncertainty [dB] | ¦Mean Error¦ [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TRP | < 3GHz | sin() | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.13 | 0 |
| TRP |  | Clenshaw-Curtis | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.10 | 0 |
| TRP (Note 2) |  |  | 30 | 30 | 62 | 0.10 | 0 |
| TRS (Note 2) |  |  | 30 | 60 | 32 | 0.15 | 0 |
| TRS (Note 2) |  |  | 45 | 45 | 26 | 0.19 | 0 |
| TRP | > 3GHz | sin() | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.21 | 0 |
| TRP |  | Clenshaw-Curtis | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.18 | 0 |
| TRP (Note 2) |  |  | 30 | 30 | 62 | 0.18 | 0 |
| TRS (Note 2) |  |  | 30 | 60 | 32 | 0.25 | 0 |
| TRS (Note 2, 3) |  |  | 45 | 45 | 27 | 0.23 | 0.08 |
| TRS (Note 2, 4) |  |  | 45 | 45 | 25 | 0.25 | 0.08 |
| Note 1: The exact number of grid points depends on how the back pole EIRP(=180°)/EIS(=180°) is approximated due to obstruction and/or blocking Note 2: The overall MU shall not be larger than the maximum MU limits if the coarsest measurement grid is adopted. Note 3: The EIS value at 180° is determined from two 165º measurements. Note 4: The EIS value at 180° is averaged from previous cut. |  |  |  |  |  |  |  |

The mean error in Table A.4.2.12-1 shall be considered a systematic uncertainty that cannot be corrected and thus shall be included in the uncertainty budget table as a systematic uncertainty added to the combined expanded uncertainty.

The legacy grid with Dq=Df=15° for TRP and Dq=Df=30° for TRS should be considered the default measurement grid and is recommended for certification testing.

<<< Skip Unchanged Sections >>>

#### A.4.3.1	Anechoic Chamber Method

The uncertainty contributions related to TRP for the Anechoic Chamber method are listed in Table A.4.3.1-1. Example uncertainty budget is presented in Table A.4.3.1-2.

Table A.4.3.1-1: Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure A.3.1-1, Figure A.3.1-2) |  |  |
| 1 | Mismatch of receiver chain | A.4.2.1 |
| 2 | Insertion loss of receiver chain | A.4.2.2 |
| 3 | Influence of the measurement antenna cable | A.4.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | A.4.2.4 |
| 5 | Measurement distance | A.4.2.7.1 |
| 6 | Quality of quiet zone | A.4.2.8 |
| 7 | DUT Tx-power drift | A.4.2.9 |
| 8 | Uncertainty related to the use of phantoms | A.4.2.11 |
| 9 | Coarse sampling grid | A.4.2.12 |
| 10 | Random uncertainty | A.4.2.13 |
| 11 | Frequency Response | A.4.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure A.3.2-1) |  |  |
| 12 | Uncertainty of network analyzer | A.4.2.15 |
| 13 | Mismatch of receiver chain | A.4.2.1 |
| 14 | Insertion loss of receiver chain | A.4.2.2 |
| 15 | Mismatch in the connection of calibration antenna | A.4.2.1 |
| 16 | Influence of the calibration antenna feed cable | A.4.2.3 |
| 17 | Influence of the measurement antenna cable | A.4.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | A.4.2.16 |
| 19 | Measurement distance | A.4.2.7.2 |
| 20 | Quality of the Quiet Zone | A.4.2.8 |
| Systematic Errors |  |  |
| 21 | Systematic Error related to TRP grids | A.4.2.12 |

Table A.4.3.1-2: Example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.41 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | From datasheet of communication tester or spectrum analyser. | 0.42 | 0.54 | Normal | 2.00 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.73 | 1 | 0.37 | 0.37 |
| 9 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0 | 0 | Actual | 1.00 | 1 | 0.0 | 0.0 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2.00 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: phase center is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  |  | 0.84 | 0.88 |  |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  | 1.64 | 1.73 |  |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.64 | 1.73 |

<<< Skip Unchanged Sections >>>

#### A.4.4.1	Anechoic Chamber Method

The uncertainty contributions related to TRS for the Anechoic Chamber method are listed in Table A.4.4.1-1. Example uncertainty budget is presented in Table A.4.4.1-2.

Table A.4.4.1-1: Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure A.3.1-1, Figure A.3.1-2) |  |  |
| 1 | Mismatch of transmitter chain | A.4.2.1 |
| 2 | Insertion loss of transmitter chain | A.4.2.2 |
| 3 | Influence of the measurement antenna cable | A.4.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | A.4.2.5 |
| 5 | Sensitivity measurement: output level step resolution | A.4.2.6 |
| 6 | Measurement distance | A.4.2.7.1 |
| 7 | Quality of quiet zone | A.4.2.8 |
| 8 | DUT sensitivity drift | A.4.2.10 |
| 9 | Uncertainty related to the use of phantoms | A.4.2.11 |
| 10 | Coarse sampling grid | A.4.2.12 |
| 11 | Random uncertainty | A.4.2.13 |
| 12 | Frequency Response | A.4.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure A.3.2-1) |  |  |
| 13 | Uncertainty of network analyzer | A.4.2.15 |
| 14 | Mismatch of transmitter chain | A.4.2.1 |
| 15 | Insertion loss of transmitter chain | A.4.2.2 |
| 16 | Mismatch in the connection of calibration antenna | A.4.2.1 |
| 17 | Influence of the calibration antenna feed cable | A.4.2.3 |
| 18 | Influence of the measurement antenna cable | A.4.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | A.4.2.16 |
| 20 | Measurement distance | A.4.2.7.2 |
| 21 | Quality of quiet zone | A.4.2.8 |
| Systematic Errors |  |  |
| 22 | Systematic Error related to TRS grids | A.4.2.12 |

Table A.4.4.1-2: Example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | Гreceiver < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.41 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | From datasheet of communication tester | 1.30 | 1.30 | Normal | 2.00 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.73 | 1 | 0 | 0 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.73 | 1 | 0.37 | 0.37 |
| 10 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0.10 | 0.18 | Actual | 1.00 | 1 | 0.10 | 0.18 |
| 11 | Random uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2.00 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of transmitter chain | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: phase center is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 1.05 | 1.09 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.06 | 2.13 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.00 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.06 | 2.13 |

<<< END OF CHANGES >>>