3GPP TSG-RAN5 Meeting #104 	R5-244801

# Maastricht, Netherlands, 19th Aug. 2024 - 23rd Aug. 2024

| CR-Form-v12.3 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0792 | rev | - | Current version: | 18.3.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Update of FR2c MU |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Anritsu |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI16_Test, NR_bands_BW_R16-UEConTest |  |  |  |  |  | Date: |  |  | 2024-08-09 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-18 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) Rel-20 (Release 20) |  |
| Reason for change: |  | - [ ] is remaining for some MU values for FR2c. - MTSU of OBW for FR2c is TBD for CBW ≥ 200 MHz. |  |  |  |  |  |  |  |  |
| Summary of change: |  | - [ ] in the following test cases was removed from MU values for FR2c. - MPR - Minimum output power - Transmit OFF power - ACLR - ACS - MTSU of OBW CBW ≥ 200 MHz were defined for FR2c. - Editorial corrections |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Test cases will not be completed for FR2c. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.2.2.27, B.4, B.7, B.8, B.15, B.17, B.21 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | x | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | x | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | x | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Depending on the outcome of MU discussion R5-244798 |  |  |  |  |  |  |  |  |
| This CR's revision history: |  |  |  |  |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.2.2.27	Influence of noise

See B.2.1.27.

The uncertainty value of influence of noise is estimated as below table and used across clause B.

Editor’s Note: For ACLR, all applicable configurations need to be added.

Table B.2.2.27-1: Uncertainty value for influence of noise for PC3, SISO for IFF

| Test case | Frequency range | Noise floor | Minimum requirement | Estimated SNRtotal [dB/400MHz] | Relaxation | Influence of noise |
| --- | --- | --- | --- | --- | --- | --- |
| MOP-EIRP | FR2a | N/A | 20.7dBm/ChBW (22.4-1.7) | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 18.9dBm/ChBW (20.6-1.7) | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2c | N/A | 18.2dBm/ChBW (18.7-0.5) | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-TRP | FR2a | N/A | 23dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b, FR2c | N/A | 23dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-Spherical | FR2a | N/A | 9.75dBm/ChBW (Spherical – MBR= 11.5-1.75) | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2b | N/A | 7.6dBm/ChBW (Spherical – MBR=8-0.4) | 6.37 (NOTE 1) | 0 | 0.9 |
|  | FR2c | N/A | 5.4dBm/ChBW (Spherical – MBR=5.8-0.4) | 5.86 (NOTE 1) | 0 | 1.0 |
| MPR | FR2a | -7.6dBm/400MHz | 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 15.17 (NOTE 1) | 0 | 0.13 |
|  | FR2b | -5.5dBm/400MHz | 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 11.30 (NOTE 1) | 0 | 0.31 |
|  | FR2c | -4.5dBm/400MHz | 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | 8.70 (NOTE 1) | 0 | 0.55 |
| Configured output power with power boost | FR2a | N/A | 21.7dBm/ChBW (22.4-1.7+1) | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 19.9dBm/ChBW (20.6-1.7+1) | 11.45 (NOTE 1) | 0 | 0.3 |
| Minimum output power | FR2a | -10.6dBm/400MHz | -13dBm | -2.54 (NOTE 1) | 8.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz | -13dBm | -7.64 (NOTE 1) | 13.5 | 1.0 (with relaxation) |
|  | FR2c | -4.5dBm/400MHz | -13dBm | -8.64 (NOTE 1) | 14.5 | 1.0 (with relaxation) |
| OFF power – TRP | FR2a | N/A | -35dBm/ChBW | -24.54 (NOTE 2) | 30.4 | 1.0 (with relaxation) |
|  | FR2b | N/A |  | <-24.54 (NOTE 2) | N/A | Propose not to test |
|  | FR2c | -4.5dBm/400MHz |  | -30.64 (NOTE 2) | 36.5 | 1.0 (with relaxation) |
| OFF power – EIRP | FR2a | -7.6dBm/400MHz | -30dBm/ChBW | -22.54 (NOTE 2) | 28.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz |  | -24.64 (NOTE 2) | 30.5 | 1.0 (with relaxation) |
|  | FR2c | [-4.5]dBm/400MHz |  | [-28.64] (NOTE 2) | [34.5] | [1.0] (with relaxation) |
| Absolute power tolerance | Same as Minimum output power |  |  |  |  |  |
| Relative power tolerance | FR2a | -13.6dBm/100MHz | -7.6dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -11.5dBm/100MHz | -5.5dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
| Aggregate power tolerance | Same as Relative power tolerance |  |  |  |  |  |
| Aggregate power tolerance | FR2a | -13.6dBm/100MHz | -7.6dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -11.5dBm/100MHz | -5.5dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
| SEM | FR2a | N/A | -13dBm/1MHz | 8.14 (NOTE 1) | 0 | 0.62 |
|  | FR2b, FR2c | N/A |  | 5.86 (NOTE 1) | 0 | 1.0 |
| ACLR (CP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB 16.65dBm/ChBW (EIRP-MPB-MPR-T(MPR) =22.4-0.75-3-2) Actual lowest: 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 22.86 (with 3dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 16.35dBm/ChBW (EIRP-MPB-MPR-T(MPR) =20.6-0.75-2-1.5) Actual lowest: 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 21.86 (with 2dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2c | -4.5dBm/400MHz | Highest testable MPR for 400MHz: 0dB 18.2dBm/ChBW (EIRP-MPB-MPR-T(MPR) =18.7-0.5-0-0) Actual lowest: 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | 22.71 (with 0dB MPR) (NOTE 1) | 0 | N/A |
| ACLR (ACP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB -0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-3-2-17) Actual lowest: -9.35 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-9-5-17) | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-2-1.5-16) Actual lowest: -10.15 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-9-5-16) | 5.86 (with 2dB MPR) (NOTE 1) | 0 | 1.0 |
|  | FR2c | -4.5dBm/400MHz | Highest testable MPR for 400MHz: 0dB 2.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=18.7-0.5-0-0-16) Actual lowest: -11.8dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=18.7-0.5-9-5-16) | 6.71 (with 0dB MPR) (NOTE 1) | 0 | 1.14 |
| General Tx spurious | 6GHz <=f<=23.45GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 | 0.41 |
|  | 23.45GHz<=f<=40GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 |  |
|  | 40GHz<=f<=80GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 |  |
|  | 80GHz<f<=87GHz | N/A | -13dBm/1MHz | 6.8 (NOTE 1) | 0 | 0.83 |
| Tx spurious Co-existence | n260 (Aggressor band : n257, n261) | -23 | -2dBm/100MHz (-22dBm/MHz) | 0.86 (NOTE 2) | 5 | 1.0 (with relaxation) |
|  | n257, n261 (Aggressor band : n260) | -27.7 | -5dBm/100MHz (-25dBm/MHz) | 2.56 (NOTE 2) | 3.3 | 1.0 (with relaxation) |
|  | 23.6 GHz ≤ f ≤ 24.0GHz | -27.7 | 1dBm/200MHz (-22 dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
|  | 36 GHz ≤ f ≤ 37GHz | -23dBm/MHz | 7dBm/1000MHz (-23dBm/MHz) | -0.14 (NOTE 2) | 6 | 1.0 (with relaxation) |
|  | 57 GHz ≤ f ≤ 66GHz | N/A | 2dBm/100MHz (-18dBm/MHz) | 5.86 (NOTE 1) | 0 | 1.0 |
| Additional spurious emission | NS_202 (7.25GHz <=f <=12.75GHz) | -40 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | 10 (NOTE 1) | 0 | 0.41 |
|  | NS_202 (12.75GHz <=f <=23.45GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_202 (23.6GHz <=f <=24.0GHz) | -27.7 dBm/MHz | 1dBm/200MHz (-22 dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
|  | NS_202 (23.45GHz <=f <=40.8GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_202 (40.8GHz <=f <=66GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_203 (23.6GHz <=f <=24.0GHz) | -27.7 dBm/MHz | +1dBm/200MHz (-22dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
| Rx spurious | 6GHz <=f<=20GHz |  | -47dBm/1MHz | -2.2 (NOTE 3) | 10.2 | 1.0 dB for 23.45~40.8GHz and 80~87 GHz, 0.64dB for 6~23.45 and 40.8~80 GHz. |
|  | 20GHz<=f<=40GHz |  | -47dBm/1MHz | -11.34 (NOTE 2) | 17.2 |  |
|  | 40GHz<=f<=80GHz |  | -47dBm/1MHz | -25.1 (NOTE 3) | 33.1 |  |
|  | 80GHz<f<=87GHz |  | -47dBm/1MHz | -27.1 (NOTE 2) | 33.1 |  |
| NOTE 1: Estimated SNR is calculated based on agreed influence of noise. NOTE 2: Estimated SNR is calculated based on agreed relaxation value: Estimated SNR = 6dB - relaxation. NOTE 3: Estimated SNR is calculated based on agreed relaxation value: Estimated SNR = 8dB - relaxation. |  |  |  |  |  |  |

Table B.2.2.27-1a: Uncertainty value for influence of noise for PC3, MIMO for IFF

| Test case | Frequency range | Noise floor | Minimum requirement | Estimated SNRtotal [dB/400MHz] | Relaxation | Influence of noise |
| --- | --- | --- | --- | --- | --- | --- |
| MOP-EIRP | FR2a | N/A | 22.4 dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 20.6 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2c | N/A | 18.7 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MPR | FR2a | -7.6dBm/400MHz | 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 15.17 (NOTE 1) | 0 | 0.13 |
|  | FR2b | -5.5dBm/400MHz | 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 11.30 (NOTE 1) | 0 | 0.31 |
|  | FR2c | -4.5dBm/400MHz | 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | 8.70 (NOTE 1) | 0 | 0.55 |
| MOP-TRP | FR2a | N/A | 23 dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b, FR2c | N/A | 23 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-Spherical | FR2a | N/A | 11.5 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2b | N/A | 8 dBm/ChBW | 6.37 (NOTE 1) | 0 | 0.9 |
| Minimum output power | FR2a | -10.6dBm/400MHz | -10 dBm | -0.46 (NOTE 1) | 5.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz | -10 dBm | -4.64 (NOTE 1) | 10.5 | 1.0 (with relaxation) |
| OFF power – TRP | FR2a | N/A | -35dBm/ChBW | -24.54 (NOTE 2) | 30.4 | 1.0 (with relaxation) |
|  | FR2b | N/A |  | <-24.54 (NOTE 2) | N/A | Propose not to test |
| SEM | FR2a | N/A | -13dBm/1MHz | 8.14 (NOTE 1) | 0 | 0.62 |
|  | FR2b, FR2c | N/A |  | 5.86 (NOTE 1) | 0 | 1.0 |
| ACLR (CP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB 16.65dBm/ChBW (EIRP-MPB-MPR-T(MPR) =22.4-0.75-3-2) Actual lowest: 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 22.86 (with 3dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 16.35dBm/ChBW (EIRP-MPB-MPR-T(MPR) =20.6-0.75-2-1.5) Actual lowest: 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 21.86 (with 2dB MPR) (NOTE 1) | 0 | N/A |
| ACLR (ACP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB -0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-3-2-17) Actual lowest: -9.35 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-9-5-17) | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-2-1.5-16) Actual lowest: -10.15 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-9-5-16) | 5.86 (with 2dB MPR) (NOTE 1) | 0 | 1.0 |
| General Tx spurious | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| Tx spurious Co-existence | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| Additional spurious emission | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| NOTE 1: Estimated SNR is calculated based on agreed influence of noise. |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.4	UE maximum output power for modulation / channel bandwidth

Following tables summarize the MU threshold for EIRP measurements for UE maximum output power for modulation / channel bandwidth (a.k.a Maximum Power Reduction/MPR). The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.4-1: MU threshold for EIRP measurement for UE maximum output power for modulation / channel bandwidth

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MBR - MPR – T(MPR) | 5.11 | 5.38 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.29 | 5.56 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.89 | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MPR – T(MPR) | 5.33 | 5.60 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.50 | 5.77 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MPR – T(MPR) | 5.33 | 5.60 |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.4.2-2 and Table B.4.2-3 for PC3 UEs and Table B.4.2-4 and Table B.4.2-5 for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. NOTE 3: The MU values are valid for SISO and MIMO. |  |  |  |  |  |

B.4.1	Uncertainty budget format and assessment for DFF

<<Unchanged sections skipped>>

B.4.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.4.2-1.

Table B.4.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Same as Stage 2 of Table 3.2-1 for EIRP |  |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, P = maximum output power – MPR – T(MPR).

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.2-2 for PC3 UEs and in Table B.4.2-x for PC1 UEs.

Table B.4.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 to 16 | Same as Stage 2 of Table B.3.2-2 for EIRP |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 to 27 | Same as Stage 1 of Table B.3.2-2 for EIRP |  |  |  |  |
|  | Systematic uncertainties |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature |  |  |  | N/A |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.31 |
| 29 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | 0.55 |
| 30 | Systematic error related to beam peak search |  |  |  | Same as Table B.3.2-2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ – confidence interval of 95 %) [dB] |  |  |  |  | 5.11 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ – confidence interval of 95 %) [dB] |  |  |  |  | 5.29 |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ – confidence interval of 95 %) [dB] |  |  |  |  | 6.89 |
| NOTE 1: The assessment assumes maximum DUT output power – MBR - MPR – T(MPR) NOTE 2: The analysis is valid for SISO and MIMO. |  |  |  |  |  |

<<Unchanged sections skipped>>

B.7	Minimum Output power

Following tables summarize the MU threshold for EIRP measurements for Minimum Output Power. The origin MU values for different test setups can be found in following clauses.

Table B.7-1: MU threshold for EIRP measurement for Minimum output power

| Power Class | Diversity | Frequency | MBW | Power | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value For ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- | --- |
| PC1 | SISO | 23.45GHz ≤ f ≤ 32.125GHz | BW ≤ 400MHz | P = Minimum Output Power | 5.66 | 5.92 |
|  |  | 32.125GHz < f ≤ 40.8GHz |  |  | 5.96 | 6.22 |
|  | MIMO | 23.45GHz ≤ f ≤ 32.125GHz |  |  | 5.51 | 5.77 |
|  |  | 32.125GHz < f ≤ 40.8GHz |  |  | 5.66 | 5.92 |
| PC3 | SISO, MIMO | 23.45GHz ≤ f ≤ 32.125GHz | BW ≤ 400MHz | P = Minimum Output Power | 6.15 | 6.41 |
|  |  | 32.125GHz < f ≤ 40.8GHz |  |  | 6.15 | 6.41 |
|  |  | 40.8GHz < f ≤ 44.3GHz |  |  | 7.34 | TBD |
| PC5 | SISO, MIMO | 23.45GHz ≤ f ≤ 32.125GHz | BW ≤ 400MHz | P = Minimum Output Power | 6.36 | 6.62 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.7.2-2 and Table B.7.2-5 for PC3 UEs and Table B.7.2-3 and Table B.7.2-6 for PC1 and PC5 UEs. |  |  |  |  |  |  |

B.7.1	Uncertainty budget format and assessment for DFF

<<Unchanged sections skipped>>

B.7.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.7.2-1.

Table B.7.2-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of beam peak search grid | B.2.2.23 |
| 14 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 15 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.2.4 |
| 17 | Amplifier Uncertainties | B.2.2.8 |
| 18 | Misalignment of positioning System | B.2.2.13 |
| 19 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 22 | Phase centre offset of calibration antenna | B.2.2.18 |
| 23 | Quality of quiet zone for calibration process | B.2.2.19 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 25 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 26 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 27 | Systematic error related to beam peak search | B.2.2.28 |
| 28 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, P = Minimum output power.

-	The uncertainty assessment for EIRP is provided in Table B.7.2-2 for PC3 UEs and in Table B.7.2-3 for PC1 UEs.

Table B.7.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 8) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (NOTE 1) (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (23.45GHz <= f <= 40.8GHz) | 2.50 | Normal | 2.00 | 1.25 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning (NOTE 3) | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 28 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1.0 |
| 28 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | 1.0 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 7.34 |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA and is valid for SISO and MIMO. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

<<Unchanged table skipped>>

NOTE: MU assessment in Table B.7.2-2 and Table B7.2-3 is based on the following relaxations for 400MHz BW:

Table B.7.2-4: Minimum output power requirement relaxation considered in MU assessment for 400 MHz EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm)

| Power Class | Diversity | Frequency Range | Relaxation (dB) |
| --- | --- | --- | --- |
| PC1 | SISO | FR2a | 0 |
|  |  | FR2b | 0 |
| PC2 | SISO | FR2a | FFS |
|  |  | FR2b | FFS |
| PC3 | SISO | FR2a | 8.4 |
|  |  | FR2b | 13.5 |
|  |  | FR2c | 14.5 |
| PC3 | MIMO | FR2a | 5.4 |
|  |  | FR2b | 10.5 |
|  |  | FR2c | FFS |
| PC4 | SISO | FR2a | FFS |
|  |  | FR2b | FFS |
| PC5 | SISO, MIMO | FR2a | 1.4 |

<<Unchanged tables skipped>>

B.8	Transmit OFF power

Following tables summarize the MU threshold for TRP measurements for Transmit OFF power. The origin MU values for different test setups can be found in following clauses.

Table B.8-1: MU threshold for TRP measurement for Transmit OFF power

| Power Class | Diversity | Frequency | MBW | Power | Threshold MU value (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Off Power | 5.67 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | N/A |
|  |  | 40.8GHz < f <= 44.3GHz |  |  | 6.86 |
| PC1 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Off Power | 5.67 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | N/A |
| PC5 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Off Power | 5.67 |
| NOTE 1: Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.8.2-2 for PC3 UEs and Table B.8.2-6 for PC1 and PC5 UEs. |  |  |  |  |  |

Table B.8-2: Void

B.8.1	Uncertainty budget format and assessment for DFF

<<Unchanged sections skipped>>

B.8.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.8.2-1.

Table B.8.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.2.4 |
| 18 | Amplifier Uncertainties | B.2.2.8 |
| 19 | Misalignment of positioning System | B.2.2.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 23 | Phase centre offset of calibration antenna | B.2.2.18 |
| 24 | Quality of quiet zone for calibration process | B.2.2.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 27 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 29 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3 GHz}, P = Off power.

-	The uncertainty assessment for TRP is provided in Table B.8.2-2 for PC3 UEs and Table B.8.2-6 for PC1 and PC5 UEs.

Table B.8.2-2: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 8) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (NOTE 1) (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (23.45GHz <= f <= 40.8GHz) | 2.50 | Normal | 2.00 | 1.25 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.5 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.7 | Normal | 2.00 | 0.85 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 8) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Quality of quiet zone for calibration process (NOTE 8) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.0 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | N/A |
| 29 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | 1.0 |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.67 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | N/A |
| TRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.86 |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA and is valid for SISO and MIMO. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.8.2-3: Void

Table B.8.2-4: Void

Table B.8.2-5: Void

<<Unchanged table skipped>>

NOTE: MU assessment in Table B.8.2-2 and Table B.8.2-6 is based on the following relaxation for 400MHz BW.

Table B.8.2-7: Transmit OFF power requirement relaxation considered in MU assessment for 400 MHz EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm)

| Power Class | Diversity | Frequency Range | Relaxation (dB) |
| --- | --- | --- | --- |
| PC1, PC2, PC3, PC4, PC5 | SISO, MIMO | FR2a | 30.4 |
|  |  | FR2b | FFS |
|  |  | FR2c | 36.5 |

<<Unchanged sections skipped>>

B.15	Occupied bandwidth

Following tables summarize the MU threshold for EIRP measurements for Occupied bandwidth. The origin MU values for different test setups can be found in following subclauses.

Table B.15-1: MU threshold for beam peak measurement for Occupied bandwidth

| Power Class | Frequency | Power | BW | Threshold MU value (NOTE1) [%CBW] |
| --- | --- | --- | --- | --- |
| PC3, PC1 | 23.45 GHz ≤ f ≤ 32.125 GHz | P = Max Output Power | 50 MHz | ±0.4 |
|  |  |  | 100 MHz | ±0.4 |
|  |  |  | 200 MHz | ±1.2 |
|  |  |  | 400 MHz | ±1.2 |
|  | 32.125 GHz < f ≤ 40.8 GHz | P = Max Output Power | 50 MHz | ±0.4 |
|  |  |  | 100 MHz | ±0.4 |
|  |  |  | 200 MHz | ±1.3 |
|  |  |  | 400 MHz | ±1.3 |
| PC3 | 40.8GHz < f <= 44.3GHz | P = Max Output Power | 50 MHz | ±0.65 |
|  |  |  | 100 MHz | ±0.65 |
|  |  |  | 200 MHz | ±1.3 |
|  |  |  | 400 MHz | ±1.3 |
| PC5 | 23.45 GHz ≤ f ≤ 32.125 GHz | P = Max Output Power | 50 MHz | ±0.4 |
|  |  |  | 100 MHz | ±0.4 |
|  |  |  | 200 MHz | ±1.2 |
|  |  |  | 400 MHz | ±1.2 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm. |  |  |  |  |

B.15.1	Uncertainty budget format and assessment for DFF

FFS

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz, 44.3 GHz}, P = [Maximum output power].

B.15.2	Uncertainty budget format and assessment for IFF

FFS

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3 GHz}, P = Maximum output power.

<<Unchanged sections skipped>>

B.17	Adjacent Channel Leakage Ratio

Editor’s Note: MU value analysis for PC1, 2 and 4 are not complete.

Following tables summarize the MU threshold for EIRP measurements for Adjacent Channel Leakage Ratio. The origin MU values for different test setups can be found in following subclauses.

Table B.17-1: MU threshold for TRP measurement for Adjacent Channel Leakage Ratio

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
|  | 40.8GHz < f <= 44.3GHz |  |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and B.17.2-3 for PC1 UEs. |  |  |  |  |

Table B.17-1B: MU threshold for EIRP measurement for Adjacent Channel Leakage Ratio

| Power Class | Frequency | CBW | Power | Threshold MU value for NTC and ETC (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | 50MHz | P = Max Output Power | 5.63 |
|  |  | 100MHz |  | 6.09 |
|  |  | 200MHz |  | 6.09 (NOTE5) |
|  |  | 400MHz |  | 6.09 (NOTE2) |
|  | 32.125GHz < f <= 40.8GHz | 50MHz | P = Max Output Power | 6.09 (NOTE7) |
|  |  | 100MHz |  | 6.09 (NOTE6) |
|  |  | 200MHz |  | 6.09 (NOTE3) |
|  |  | 400MHz |  | 6.09 (NOTE4) |
|  | 40.8GHz < f <= 44.3GHz | 50MHz | P = Max Output Power | 7.75 (NOTE 9) |
|  |  | 100MHz |  | 7.75 (NOTE 10) |
|  |  | 200MHz |  | 7.75 (NOTE 11) |
|  |  | 400MHz |  | 7.75 (NOTE 12) |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 6.04 |
|  | 32.125GHz < f <= 40.8GHz | BW <= 400MHz | P = Max Output Power | 6.04 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 6.04 |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and Table B.17.2-3 for PC1 and PC5 UEs. NOTE 2: This value is based on the relaxation of (MPR + T(MPR) – 6.0) dB for MPR > 3.0dB. NOTE 3: Not applicable for MPR > 3.5dB NOTE 4: Not applicable for MPR > 2.0dB NOTE 5: This value is based on the relaxation of (MPR + T(MPR) – 9.0) dB for MPR > 5.0dB. NOTE 6: Not applicable for MPR > 5.0dB NOTE 7: Not applicable for MPR > 7.5dB NOTE 8: The MU values are valid for SISO and MIMO. NOTE 9: Not applicable for MPR > 5.5dB NOTE 10: Not applicable for MPR > 4.0dB NOTE 11: Not applicable for MPR > 2.5dB NOTE 12: Not applicable for MPR > 0.0dB |  |  |  |  |

B.17.1	Uncertainty budget format and assessment for DFF

<<Unchanged sections skipped>>

B.17.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.17.2-1.

<<Unchanged tables skipped>>

Table B.17.2-4: Influence of noise measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

|  | FR2a | FR2b | FR2c |
| --- | --- | --- | --- |
| ChBW (50MHz) | 0.54 | 1.0 (NOTE 6) | 1.14 (NOTE 8) |
| ChBW (100MHz) | 1.0 | 1.0 (NOTE 5) | 1.14 (NOTE 9) |
| ChBW (200MHz) | 1.0 (NOTE 4) | 1.0 (NOTE 2) | 1.14 (NOTE 10) |
| ChBW (400MHz) | 1.0 (NOTE 1) | 1.0 (NOTE 3) | 1.14 (NOTE 11) |
| NOTE 1: This value is based on the relaxation of (MPR + T(MPR) – 6.0) dB for MPR > 3.0dB. NOTE 2: Not applicable for MPR > 3.5dB NOTE 3: Not applicable for MPR > 2.0dB NOTE 4: This value is based on the relaxation of (MPR + T(MPR) – 9.0) dB for MPR > 5.0dB. NOTE 5: Not applicable for MPR > 5.0dB NOTE 6: Not applicable for MPR > 7.5dB NOTE 7: Values are valid for SISO and MIMO. NOTE 8: Not applicable for MPR > 5.5dB NOTE 9: Not applicable for MPR > 4.0dB NOTE 10: Not applicable for MPR > 2.5dB NOTE 11: Not applicable for MPR > 0.0dB |  |  |  |

<<Unchanged sections skipped>>

B.21	Adjacent Channel Selectivity

Following tables summarize the MU threshold for Adjacent Channel Selectivity measurement. The origin MU values for different test setups with varies parameters can be found in following subclauses.

Table B.21-1: MU threshold for Adjacent Channel Selectivity

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 8.08 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 8.08 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 9.46 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 8.31 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 8.31 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 8.31 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.21.2-2 for PC3 UEs and Table B.21.2-3 for PC1 UEs. |  |  |  |  |

B.21.1	Uncertainty budget format and assessment for DFF

FFS

B.21.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.21.2-1.

Table B.21.2-1: Total Uncertainty contributions for Adjacent Channel Selectivity measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |
| 1 | Positioning misalignment | B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | gNB emulator uncertainty | B.2.2.17 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 14 | DUT repositioning | B.2.2.26 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |
| 15 | Positioning misalignment | B.2.2.1 |
| 16 | Measure distance uncertainty | B.2.2.2 |
| 17 | Quality of Quiet Zone | B.2.2.3 |
| 18 | Mismatch | B.2.2.4 |
| 19 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 20 | Modulated Interferer uncertainty | B.2.2.33 |
| 21 | Phase curvature | B.2.2.7 |
| 22 | Amplifier uncertainties | B.2.2.8 |
| 23 | Random uncertainty | B.2.2.9 |
| 24 | Influence of the XPD | B.2.2.10 |
| 25 | Insertion Loss Variation | B.2.2.11 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 27 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 28 | DUT repositioning | B.2.2.26 |
| 29 | Influence of offset antenna (Std.Dev) | B.2.2.35 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |
| 30 | Mismatch | B.2.2.4 |
| 31 | Amplifier Uncertainties | B.2.2.8 |
| 32 | Misalignment of positioning System | B.2.2.13 |
| 33 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 36 | Phase centre offset of calibration antenna | B.2.2.18 |
| 37 | Quality of quiet zone for calibration process | B.2.2.19 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 39 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 40 | Insertion Loss Variation | B.2.2.11 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |
| 41 | Mismatch | B.2.2.4 |
| 42 | Amplifier Uncertainties | B.2.2.8 |
| 43 | Misalignment of positioning System | B.2.2.13 |
| 44 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 47 | Phase centre offset of calibration antenna | B.2.2.18 |
| 48 | Quality of quiet zone for calibration process | B.2.2.19 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 50 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 51 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 52 | Systematic error related to beam peak search | B.2.2.28 |
| 53 | Additional impact of interferer ACLR | B.2.2.32 |
| 54 | Influence of offset antenna (mean error) | B.2.2.35 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}

-	The uncertainty assessment for ACS is provided in Table B.21.2-2 for PC3 UEs and Table B.21.2-3 for PC1 and PC5 UEs.

Table B.21.2-2: Uncertainty assessment for Adjacent Channel Selectivity measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level (23.45GHz <= f <= 40.8GHz) | 2.9 | Normal | 2.00 | 1.45 |
| 6 | gNB uncertainty on absolute level (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |  |  |  |
| 15 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 16 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of Quiet Zone (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 17 | Quality of Quiet Zone (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 18 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 18 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 19 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 20 | Modulated Interferer uncertainty on absolute level (23.45GHz <= f <= 40.8GHz) | 2.9 | Normal | 2.00 | 1.45 |
| 20 | Modulated Interferer uncertainty on absolute level (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 21 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 22 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 23 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 24 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 27 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 28 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| 29 | Influence of offset antenna (Std.Dev) (NOTE 5) | 0.00 | Normal | 2.00 | 0.00 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |  |  |  |
| 30 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 31 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 32 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 33 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 33 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 36 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 37 | Quality of quiet zone for calibration process (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 37 | Quality of quiet zone for calibration process (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 39 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 40 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |  |  |  |
| 41 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 42 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 43 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 44 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 44 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 47 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 48 | Quality of quiet zone for calibration process (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 48 | Quality of quiet zone for calibration process (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 50 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 51 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 52 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 53 | Additional impact of interferer ACLR |  |  |  | 0.7 |
| 54 | Influence of offset antenna (mean error) (NOTE 5) |  |  |  | 0.00 |
| Total measurement uncertainty |  |  |  |  | Value |
| ACS Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 8.08 |
| ACS Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 9.46 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 5: For MTSU derivation purpose, this value is set to 0.0 (no offset antenna case). |  |  |  |  |  |

<<Unchanged table skipped>>

NOTE: MU assessment in Table B.21.2-2 and Table B.21.2-3 is based on the following relaxation for 200MHz BW.

Table B.21.2-4: Adjacent channel selectivity requirement relaxation considered in MU assessment for 200 MHz EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm)

| Power Class | Frequency Range | Relaxation (dB) |
| --- | --- | --- |
| PC3 | FR2a | 0 |
|  | FR2b | 4.8 |
|  | FR2c | 6.8 |
| PC1 | FR2a | 0 |
|  | FR2b | 0 |
| PC5 | FR2a | 0 |

<<End of change>>