3GPP TSG-RAN5 Meeting #105	R5-247858

# Orlando, US, 18th Nov 2024 – 22nd Nov 2024

| CR-Form-v12.3 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0860 | rev | 1 | Current version: | 18.4.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | PC7 FR2 MU definition in 38.903 |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Keysight Technologies |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI17_Test, NR_redcap_plus_ARCH-UEConTest |  |  |  |  |  | Date: |  |  | 2024-11-07 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-18 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) Rel-20 (Release 20) |  |
| Reason for change: |  | Progress has been made in the MU analysis for PC7 in discussions R5-244630 and R5-246537. In summary, for FR2a and CBW up to 100MHz, all MU factors for PC7 can be defined same as PC3 and so the MTSU calcuation for most of the test cases. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Defined MU factor and MTSU for PC7 in FR2a SISO up to 100MHz CBW. For MTSU, only considered test cases listed in the associated work plan. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | MU for PC7 will remain undefined and PC7 untestable. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.2.2 (all subclauses), B.3, B.4, B.7, B.11, B.19 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Influence of noise for Carrier leakage for PC3 was not captured in section B.2.2.27 but it is being added by R5-246245. |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | Revision 1: -Tagged PC7 in Table B.2.2.28-1, that was missing. |  |  |  |  |  |  |  |  |

## <<< START OF CHANGES >>>

## B.2.2	Measurement error contribution descriptions for IFF

### B.2.2.1	Positioning misalignment

See B.2.1.1.

The uncertainty value of positioning misalignment is estimated as below table and used across clause B.

Table B.2.2.1-1: Uncertainty value for positioning misalignment for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC5 | 0.02 | Normal | 2.00 | 0.01 |
| PC3, PC6, PC7 | 0.00 | Normal | 2.00 | 0.00 |

### B.2.2.2	Measure distance uncertainty

See B.2.1.2. For IFF1 this can be considered to be zero.

The uncertainty value of measure distance uncertainty is estimated as below table and used across clause B.

Table B.2.2.2-1: Uncertainty value for measure distance uncertainty for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Rectangular | 1.73 | 0.00 |

### B.2.2.3	Quality of Quiet Zone

See B.2.1.3.

The uncertainty value of quality of quiet zone is estimated as below table and used across clause B.

Table B.2.2.3-1: Uncertainty value for quality of quiet zone for IFF

| QZ size | Power class | Condition | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| All | All | All | ACLR (relative measurement) | FR2a, FR2b, FR2c | 0.52 | Actual | 1.00 | 0.52 |
| ≤30cm | PC1, PC3, PC5, PC6, PC7 | NC | NOTE1 | FR2a, FR2b | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | FR2c, FR2d | 0.7 | Actual | 1.00 | 0.7 |
|  |  |  | SE | 6GHz to 12.75GHz | 0.7 | Actual | 1.00 | 0.7 |
|  |  |  |  | 12.75GHz to 23.45GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 23.45GHz to 40.8GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 40.8GHz to 66GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 66GHz to 80GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 80GHz to 87GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  | ETC | NOTE1 | FR2a, FR2b, FR2c, FR2d | 0.9 | Actual | 1.00 | 0.9 |
| 40cm | PC1, PC3, PC5, PC6, PC7 | NC | NOTE1 | FR2a, FR2b | 0.7 (NOTE 2) | Actual | 1.00 | 0.7 |
|  |  |  |  | FR2c, FR2d | 0.8 (NOTE 2) | Actual | 1.00 | 0.8 |
|  |  |  | SE | 6GHz to 12.75GHz | 1.0 (NOTE 4) | Actual | 1.00 | 1.0 |
|  |  |  |  | 12.75GHz to 23.45GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 23.45GHz to 40.8GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 40.8GHz to 66GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 66GHz to 80GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 80GHz to 87GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  | ETC | NOTE1 | FR2a, FR2b | TBD | Actual | 1.00 | TBD |
| NOTE 1: The uncertainty in current row applies to maximum output power with EIRP and TRP, EIRP spherical coverage, MPR, configured output power with power boost, minimum output power, transmit OFF power, spectrum emission mask, reference sensitivity, adjacent selectivity, in-band blocking. NOTE 2: For Uncertainty assessment in tables across clause B and MTSU calculation, QoQZ uncertainty value for QZ ≤ 30cm is used also for 40cm for FR2a through FR2c. For uncertainty assessment in tables across clause B and MTSU calculation for FR2d, QoQZ uncertainty value for QZ size of 40cm is used also for QZ size ≤ 30cm. NOTE 3: Void NOTE 4: For Uncertainty assessment in tables across clause B and MTSU calculation, QoQZ uncertainty value for QZ ≤ 30cm is used also for 40cm. |  |  |  |  |  |  |  |  |

### B.2.2.4	Mismatch

See B.2.1.4.

The uncertainty value of mismatch is estimated as below table and used across clause B.

Table B.2.2.4-2: Uncertainty value for mismatch for IFF

| QZ size | Power class | Condition | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| <= 30cm | PC1, PC3, PC5, PC6, PC7 | NC | Default (FR2a, FR2b) | 1.30 | Actual | 1.00 | 1.30 |
|  |  |  | Default (FR2c) | 1.70 | Actual | 1.00 | 1.70 |
|  |  |  | ACLR (relative measurement) (FR2a, FR2b) | 1.84 | Actual | 1.00 | 1.84 |
|  |  |  | ACLR (relative measurement) (FR2c) | 2.4 | Actual | 1.00 | 2.4 |
|  |  |  | Tx SE (6GHz to 12.75GHz) | 1.5 | Actual | 1.00 | 1.5 |
|  |  |  | Tx SE (12.75GHz to 23.45GHz) | 1.5 | Actual | 1.00 | 1.5 |
|  |  |  | Tx SE (23.45GHz to 40.8GHz) | 1.4 | Actual | 1.00 | 1.4 |
|  |  |  | Tx SE (40.8GHz to 66GHz) | 2.3 | Actual | 1.00 | 2.3 |
|  |  |  | Tx SE (66GHz to 87GHz) | 2.3 | Actual | 1.00 | 2.3 |
|  |  |  | Rx SE (6GHz to 12.75GHz) | 1.6 | Actual | 1.00 | 1.6 |
|  |  |  | Rx SE (12.75GHz to 23.45GHz) | 1.6 | Actual | 1.00 | 1.6 |
|  |  |  | Rx SE (23.45GHz to 40.8GHz) | 1.5 | Actual | 1.00 | 1.5 |
|  |  |  | Rx SE (40.8GHz to 66GHz) | 2.3 | Actual | 1.00 | 2.3 |
|  |  |  | Rx SE (66GHz to 87GHz) | 2.3 | Actual | 1.00 | 2.3 |
|  |  | ETC | Default (FR2a, FR2b) | 1.30 | Actual | 1.00 | 1.30 |
|  |  |  | Default (FR2c) | 1.70 | Actual | 1.00 | 1.70 |
|  |  |  | ACLR (relative measurement) (FR2a, FR2b) | 1.84 | Actual | 1.00 | 1.84 |
|  |  |  | ACLR (relative measurement) (FR2c) | 2.4 | Actual | 1.00 | 2.4 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| <= 30cm | PC1, PC3, PC5, PC6, PC7 | NC | All | 0.00 | U-shaped | 1.41 | 0.00 |
|  |  | ETC | All | 0.00 | U-shaped | 1.41 | 0.00 |

### B.2.2.5	Standing wave between DUT and measurement antenna

See B.2.1.5.

The uncertainty value of standing wave between the DUT and measurement antenna is estimated as below table and used across clause B.

Table B.2.2.5-1: Uncertainty value for standing wave between the DUT and measurement antenna for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | U-shaped | 1.41 | 0.00 |

### B.2.2.6	Uncertainty of the RF power measurement equipment

See B.2.1.6.

The uncertainty value of RF power measurement equipment is estimated as below table and used across clause B.

Table B.2.2.6-1: Uncertainty value for RF power measurement equipment for IFF

| Power class | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | MOP, MPR, Configured output power with power boost,SEM, ACLR | FR2a, FR2b | 2.16 | Normal | 2.00 | 1.08 |
|  |  | FR2c | 3.60 | Normal | 2.00 | 1.80 |
|  | Minimum output power, OFF power | FR2a, FR2b | 2.50 | Normal | 2.00 | 1.25 |
|  |  | FR2c | 3.60 | Normal | 2.00 | 1.80 |
|  | SE | 6GHz to 12.75GHz | 2.00 | Normal | 2.00 | 1.00 |
|  |  | 12.75GHz to 23.45GHz | 2.16 | Normal | 2.00 | 1.08 |
|  |  | 23.45GHz to 40.8GHz | 2.73 | Normal | 2.00 | 1.37 |
|  |  | 40.8GHz to 66GHz | 4.00 | Normal | 2.00 | 2.00 |
|  |  | 66GHz to 87GHz | 4.00 | Normal | 2.00 | 2.00 |

### B.2.2.7	Phase Curvature

See B.2.1.7. For IFF1 this can be considered to be zero.

The uncertainty value of phase curvature is estimated as below table and used across clause B.

Table B.2.2.7-1: Uncertainty value for phase curvature for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | U-shaped | 1.41 | 0.00 |

### B.2.2.8	Amplifier Uncertainties

See B.2.1.8.

The uncertainty value of amplifier uncertainties is estimated as below table and used across clause B.

Table B.2.2.8-1: Uncertainty value for amplifier uncertainties for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| PC1, PC5 | Default | 2.10 | Normal | 2.00 | 1.05 |
|  | Relative power tolerance | TBD | Rectangular | 1.73 | TBD |
|  | SE (66GHz to 80GHz) | 3.0 | Normal | 2.00 | 1.50 |
| PC3, PC7 | Default | 2.10 | Normal | 2.00 | 1.05 |
|  | Relative power tolerance | 0.5 | Rectangular | 1.73 | 0.29 |
|  | SE (66GHz to 87GHz) | 3.0 | Normal | 2.00 | 1.50 |
| PC6 | Default | 2.10 | Normal | 2.00 | 1.05 |
|  | Relative power tolerance | TBD | Rectangular | 1.73 | TBD |
|  | SE (66GHz to 80GHz) | 3.0 | Normal | 2.00 | 1.50 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| PC1, PC3, PC5, PC6, PC7 | Default | 0.00 | Normal | 2.00 | 0.00 |

### B.2.2.9	Random uncertainty

See B.2.1.9.

The uncertainty value of random uncertainty is estimated as below table and used across clause B.

Table B.2.2.9-1: Uncertainty value for random uncertainty for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.5 | Normal | 2.00 | 0.25 |

### B.2.2.10	Influence of XPD

See B.2.1.10.

The uncertainty value of influence of the XPD is estimated as below table and used across clause B.

Table B.2.2.10-2: Uncertainty value for influence of the XPD for IFF

| QZ size | Power class | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| All | All | ACLR | FR2a, FR2b, FR2c | 0.00 | U-shaped | 1.41 | 0.00 |
| ≤30cm | PC1, PC3, PC5, PC6, PC7 | Default | FR2a, FR2b, FR2c | 0.01 | U-shaped | 1.41 | 0.00 |
|  |  | SE | 6GHz to 12.75GHz | 0.09 | U-shaped | 1.41 | 0.064 |
|  |  |  | 12.75GHz to 23.45GHz | 0.09 | U-shaped | 1.41 | 0.064 |
|  |  |  | 23.45GHz to 40.8GHz | 0.01 | U-shaped | 1.41 | 0.00 |
|  |  |  | 40.8GHz to 66GHz | 0.09 | U-shaped | 1.41 | 0.064 |
|  |  |  | 66GHz to 87GHz | 0.09 | U-shaped | 1.41 | 0.064 |
| 40cm | PC1, PC3, PC5, PC6, PC7 | Default | FR2a, FR2b, FR2c | 0.07 (NOTE 1) | U-shaped | 1.41 | 0.02 |
|  |  | SE | 6GHz to 12.75GHz | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 12.75GHz to 23.45GHz) | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 23.45GHz to 40.8GHz | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 40.8GHz to 66GHz | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 66GHz to 87GHz | TBD | U-shaped | 1.41 | TBD |
| NOTE 1: For Uncertainty assessment in tables across clause B and MTSU calculation, influence of the XPD uncertainty value for QZ ≤ 30cm is used also for 40cm. |  |  |  |  |  |  |  |

### B.2.2.11	Insertion Loss Variation

See B.2.1.11.

The uncertainty value of insertion loss variantion is estimated as below table and used across clause B.

Table B.2.2.11-1: Uncertainty value for insertion loss variantion for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Rectangular | 1.73 | 0.00 |

### B.2.2.12	RF leakage (from measurement antenna to receiver/transmitter)

See B.2.1.12.

The uncertainty value of RF leakage is estimated as below table and used across clause B.

Table B.2.2.12-1: Uncertainty value for RF leakage for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Actual | 1.00 | 0.00 |

### B.2.2.13	Misalignment of positioning system

See B.2.1.13.

The uncertainty value of misalignment of positioning system is estimated as below table and used across clause B.

Table B.2.2.13-1: Uncertainty value for misalignment of positioning system for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Normal | 2.00 | 0.00 |

### B.2.2.14	Uncertainty of the Network Analyzer

See B.2.1.14.

The uncertainty value of uncertainty of the network analyzer is estimated as below table and used across clause B.

Table B.2.2.14-1: Uncertainty value for uncertainty of the network analyser for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| All | Default (6GHz to 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
|  | SE (40.8GHz to 87GHz) | 1.70 | Normal | 2.00 | 0.85 |

### B.2.2.15	Uncertainty of the absolute gain of the calibration antenna

See B.2.1.15.

The uncertainty value of uncertainty of the absolute gain of the calibration antenna is estimated as below table and used across clause B.

Table B.2.2.15-1: Uncertainty value for uncertainty of the absolute gain of the calibration antenna for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | Default | 0.60 | Normal | 2.00 | 0.30 |
|  | SE (40.8GHz to 66GHz) | 1.70 | Normal | 2.00 | 0.85 |
|  | SE (66GHz to 87GHz) | 1.70 | Normal | 2.00 | 0.85 |

### B.2.2.16	Positioning and pointing misalignment between the reference antenna and the measurement antenna

See B.2.1.16.

The uncertainty value of positioning and pointing misalignment between the reference antenna and the measurement antenna is estimated as below table and used across clause B.

Table B.2.2.16-1: Uncertainty value for positioning and pointing misalignment between the reference antenna and the measurement antenna for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | Default | 0.01 | Rectangular | 1.73 | 0.00 |
|  | ACLR | 0.00 | Rectangular | 1.73 | 0.00 |
|  | SE | 0.05 | Rectangular | 1.73 | 0.03 |

### B.2.2.17	gNB emulator uncertainty

See B.2.1.17.

The uncertainty value of gNB emulator uncertainty is estimated as below table and used across clause B.

Table B.2.2.17-1: Uncertainty value for gNB emulator uncertainty for IFF

| Power class | Uncertainty value | Frequency range | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |
|  | 3.6 | FR2c | Normal | 2.00 | 1.8 |

### B.2.2.18	Phase centre offset of calibration

See B.2.1.18. For IFF1 this can be considered to be zero.

The uncertainty value of phase centre offset of calibration is estimated as below table and used across clause B.

Table B.2.2.18-1: Uncertainty value for phase centre offset of calibration for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Rectangular | 1.73 | 0.00 |

### B.2.2.19	Quality of the Quiet Zone for Calibration Process

See B.2.1.19.

The uncertainty value of quality of quiet zone for calibration process is estimated as below table and used across clause B.

Table B.2.2.19-1: Uncertainty value for quiet zone for calibration process for IFF

| QZ size | Power class | Condition | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| All | All | All | ACLR (relative measurement) | FR2a, FR2b, FR2c | 0.32 | Actual | 1.00 | 0.32 |
| ≤ 30cm | PC1, PC3, PC5, PC6, PC7 | NC | NOTE1 | FR2a, FR2b | 0.4 | Actual | 1.00 | 0.4 |
|  |  |  |  | FR2c, FR2d | 0.5 | Actual | 1.00 | 0.5 |
|  |  |  | SE | 6GHz to 12.75GHz | 0.7 | Actual | 1.00 | 0.7 |
|  |  |  |  | 12.75GHz to 23.45GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 23.45GHz to 40.8GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 40.8GHz to 66GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 66GHz to 80GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 80GHz to 87GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  | ETC | NOTE1 | FR2a, FR2b, FR2c, FR2d | 0.6 | Actual | 1.00 | 0.6 |
| 40cm | PC1, PC3, PC5, PC6, PC7 | NC | NOTE1 | FR2a, FR2b | 0.4 | Actual | 1.00 | 0.4 |
|  |  |  |  | FR2c, FR2d | 0.5 | Actual | 1.00 | 0.5 |
|  |  |  | SE | 6GHz to 12.75GHz | 1.0 (Note 3) | Actual | 1.00 | 1.0 |
|  |  |  |  | 12.75GHz to 23.45GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 23.45GHz to 40.8GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 40.8GHz to 66GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 66GHz to 80GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 80GHz to 87GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  | ETC | NOTE 1 | FR2a, FR2b, FR2c, FR2d | TBD | Actual | 1.00 | TBD |
| NOTE 1: The uncertainty in current row applies to maximum output power with EIRP and TRP, EIRP spherical coverage, MPR, configured output power with power boost, minimum output power, transmit OFF power, spectrum emission mask, reference sensitivity, adjacent selectivity, in-band blocking. NOTE 2: Void NOTE 3: For Uncertainty assessment in tables across clause B and MTSU calculation, QoQZ uncertainty value for QZ ≤ 30cm is used also for 40cm. |  |  |  |  |  |  |  |  |

### B.2.2.20	Standing wave between reference calibration antenna and measurement antenna

See B.2.1.20.

The uncertainty value of standing wave between reference calibration antenna and measurement antenna is estimated as below table and used across clause B.

Table B.2.2.20-1: Uncertainty value for standing wave between reference calibration antenna and measurement antenna for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | U-shaped | 1.41 | 0.00 |

### B.2.2.21	Influence of the calibration antenna feed cable (Flexing cables, adapters, attenuators, connector repeatability)

See B.2.1.21.

The uncertainty value of influence of the calibration antenna feed cable is estimated as below table and used across clause B.

Table B.2.2.21-1: Uncertainty value for influence of the calibration antenna feed cable for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | Default (FR2a,FR2b,FR2c) | 0.14 | Normal | 2.00 | 0.07 |
|  | SE (40.8GHz to 66GHz) | 0.28 | Normal | 2.00 | 0.14 |
|  | SE (66GHz to 87GHz) | 0.28 | Normal | 2.00 | 0.14 |

### B.2.2.22	Influence of TRP measurement grid

See B.2.1.22.

The uncertainty value of influence of TRP measurement grid is estimated as below table and used across clause B.

Table B.2.2.22-1: Uncertainty value for influence of TRP measurement grid for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC5 (Note 1) | Default | 0.25 | Actual | 1.00 | 0.25 |
|  | SE | 0.25 | Actual | 1.00 | 0.25 |
| PC3, PC7 | Default | 0.25 | Actual | 1.00 | 0.25 |
|  | SE | 0.32 | Actual | 1.00 | 0.32 |
| PC6 | Default | 0.25 | Actual | 1.00 | 0.25 |
|  | SE | 0.25 | Actual | 1.00 | 0.25 |
| Note 1: A finer measurement grid for SE (when compared to default) is required near the harmonic region, i.e., a constant-step size grid with 2522 unique grid points (Dq=Df=5°) or a constant-density grid with 1600 grid points. |  |  |  |  |  |

### B.2.2.23	Influence of beam peak search grid

See B.2.1.23.

The uncertainty value of influence of beam peak search grid is estimated as below table and used across clause B.

Table B.2.2.23-1: Uncertainty value for influence of beam peak search grid for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 | Actual | 1.00 | 0.00 |

### B.2.2.24	Systematic error due to TRP calculation/quadrature

See B.2.1.24.

The uncertainty value of systematic error due to TRP calculation/quadrature is estimated as below table and used across clause B.

Table B.2.2.24-1: Uncertainty value for systematic error due to TRP calculation/quadrature for IFF

| Power class | Uncertainty value |
| --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.00 |

### B.2.2.25	Multiple measurement antenna uncertainty

See B.2.1.25.

The uncertainty value of multiple measurement antenna uncertainty is estimated as below table and used across clause B.

Table B.2.2.25-1: Uncertainty value for multiple measurement antenna uncertainty for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.15 | Actual | 1.00 | 0.15 |

### B.2.2.26	DUT repositioning

See B.2.1.26.

The uncertainty value of DUT repositioning is estimated as below table and used across clause B.

Table B.2.2.26-1: Uncertainty value for DUT repositioning for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC5 | TRP, spherical coverage | 0.00 | Rectangular | 1.73 | 0.00 |
|  | EIRP, EIS | 0.35 | Rectangular | 1.73 | 0.20 |
| PC3, PC6, PC7 | TRP, spherical coverage | 0.00 | Rectangular | 1.73 | 0.00 |
|  | EIRP, EIS | 0.08 | Rectangular | 1.73 | 0.05 |

### B.2.2.27	Influence of noise

See B.2.1.27.

The uncertainty value of influence of noise is estimated as below table and used across clause B.

Editor’s Note: For ACLR, all applicable configurations need to be added.

Table B.2.2.27-1: Uncertainty value for influence of noise for PC3, SISO for IFF

| Test case | Frequency range | Noise floor | Minimum requirement | Estimated SNRtotal [dB/400MHz] | Relaxation | Influence of noise |
| --- | --- | --- | --- | --- | --- | --- |
| MOP-EIRP | FR2a | N/A | 20.7dBm/ChBW (22.4-1.7) | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 18.9dBm/ChBW (20.6-1.7) | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2c | N/A | 18.2dBm/ChBW (18.7-0.5) | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-TRP | FR2a | N/A | 23dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b, FR2c | N/A | 23dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-Spherical | FR2a | N/A | 9.75dBm/ChBW (Spherical – MBR= 11.5-1.75) | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2b | N/A | 7.6dBm/ChBW (Spherical – MBR=8-0.4) | 6.37 (NOTE 1) | 0 | 0.9 |
|  | FR2c | N/A | 5.4dBm/ChBW (Spherical – MBR=5.8-0.4) | 5.86 (NOTE 1) | 0 | 1.0 |
| MPR | FR2a | -7.6dBm/400MHz | 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 15.17 (NOTE 1) | 0 | 0.13 |
|  | FR2b | -5.5dBm/400MHz | 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 11.30 (NOTE 1) | 0 | 0.31 |
|  | FR2c | -4.5dBm/400MHz | 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | 8.70 (NOTE 1) | 0 | 0.55 |
| Configured output power with power boost | FR2a | N/A | 21.7dBm/ChBW (22.4-1.7+1) | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 19.9dBm/ChBW (20.6-1.7+1) | 11.45 (NOTE 1) | 0 | 0.3 |
| Minimum output power | FR2a | -10.6dBm/400MHz | -13dBm | -2.54 (NOTE 1) | 8.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz | -13dBm | -7.64 (NOTE 1) | 13.5 | 1.0 (with relaxation) |
|  | FR2c | -4.5dBm/400MHz | -13dBm | -8.64 (NOTE 1) | 14.5 | 1.0 (with relaxation) |
| OFF power – TRP | FR2a | -10.6dBm/400MHz | -35dBm/ChBW | -24.54 (NOTE 2) | 30.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz |  | -29.64 (NOTE 2) | 35.5 | 1.0 (with relaxation) |
|  | FR2c | -4.5dBm/400MHz |  | -30.64 (NOTE 2) | 36.5 | 1.0 (with relaxation) |
| OFF power – EIRP | FR2a | -7.6dBm/400MHz | -30dBm/ChBW | -22.54 (NOTE 2) | 28.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz |  | -24.64 (NOTE 2) | 30.5 | 1.0 (with relaxation) |
|  | FR2c | [-4.5]dBm/400MHz |  | [-25.64] (NOTE 2) | [31.5] | [1.0] (with relaxation) |
| Absolute power tolerance | Same as Minimum output power |  |  |  |  |  |
| Relative power tolerance | FR2a | -13.6dBm/100MHz | -7.6dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -11.5dBm/100MHz | -5.5dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
| Aggregate power tolerance | Same as Relative power tolerance |  |  |  |  |  |
| Aggregate power tolerance | FR2a | -13.6dBm/100MHz | -7.6dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -11.5dBm/100MHz | -5.5dBm/100MHz | 5.86 (NOTE 1) | 0 | 1.0 |
| SEM | FR2a | N/A | -13dBm/1MHz | 8.14 (NOTE 1) | 0 | 0.62 |
|  | FR2b, FR2c | N/A |  | 5.86 (NOTE 1) | 0 | 1.0 |
| ACLR (CP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB 16.65dBm/ChBW (EIRP-MPB-MPR-T(MPR) =22.4-0.75-3-2) Actual lowest: 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 22.86 (with 3dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 16.35dBm/ChBW (EIRP-MPB-MPR-T(MPR) =20.6-0.75-2-1.5) Actual lowest: 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 21.86 (with 2dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2c | -4.5dBm/400MHz | Highest testable MPR for 400MHz: 0dB 18.2dBm/ChBW (EIRP-MPB-MPR-T(MPR) =18.7-0.5-0-0) Actual lowest: 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | 22.71 (with 0dB MPR) (NOTE 1) | 0 | N/A |
| ACLR (ACP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB -0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-3-2-17) Actual lowest: -9.35 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-9-5-17) | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-2-1.5-16) Actual lowest: -10.15 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-9-5-16) | 5.86 (with 2dB MPR) (NOTE 1) | 0 | 1.0 |
|  | FR2c | -4.5dBm/400MHz | Highest testable MPR for 400MHz: 0dB 2.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=18.7-0.5-0-0-16) Actual lowest: -11.8dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=18.7-0.5-9-5-16) | 6.71 (with 0dB MPR) (NOTE 1) | 0 | 1.14 |
| General Tx spurious | 6GHz <=f<=23.45GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 | 0.41 |
|  | 23.45GHz<=f<=40GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 |  |
|  | 40GHz<=f<=80GHz | N/A | -13dBm/1MHz | 10.0 (NOTE 1) | 0 |  |
|  | 80GHz<f<=87GHz | N/A | -13dBm/1MHz | 6.8 (NOTE 1) | 0 | 0.83 |
| Tx spurious Co-existence | n260 (Aggressor band : n257, n261) | -23 | -2dBm/100MHz (-22dBm/MHz) | 0.86 (NOTE 2) | 5 | 1.0 (with relaxation) |
|  | n257, n261 (Aggressor band : n260) | -27.7 | -5dBm/100MHz (-25dBm/MHz) | 2.56 (NOTE 2) | 3.3 | 1.0 (with relaxation) |
|  | 23.6 GHz ≤ f ≤ 24.0GHz | -27.7 | 1dBm/200MHz (-22 dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
|  | 36 GHz ≤ f ≤ 37GHz | -23dBm/MHz | 7dBm/1000MHz (-23dBm/MHz) | -0.14 (NOTE 2) | 6 | 1.0 (with relaxation) |
|  | 57 GHz ≤ f ≤ 66GHz | N/A | 2dBm/100MHz (-18dBm/MHz) | 5.86 (NOTE 1) | 0 | 1.0 |
| Additional spurious emission | NS_202 (7.25GHz <=f <=12.75GHz) | -40 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | 10 (NOTE 1) | 0 | 0.41 |
|  | NS_202 (12.75GHz <=f <=23.45GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_202 (23.6GHz <=f <=24.0GHz) | -27.7 dBm/MHz | 1dBm/200MHz (-22 dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
|  | NS_202 (23.45GHz <=f <=40.8GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_202 (40.8GHz <=f <=66GHz) | -23 dBm/MHz | -10dBm/100MHz (-30 dBm/MHz) | -7.14 (NOTE 2) | 13 | 1.0 (with relaxation) |
|  | NS_203 (23.6GHz <=f <=24.0GHz) | -27.7 dBm/MHz | +1dBm/200MHz (-22dBm/MHz) | 5.56 (NOTE 2) | 0.3 | 1.0 (with relaxation) |
| Rx spurious | 6GHz <=f<=20GHz |  | -47dBm/1MHz | -2.2 (NOTE 3) | 10.2 | 1.0 dB for 23.45~40.8GHz and 80~87 GHz, 0.64dB for 6~23.45 and 40.8~80 GHz. |
|  | 20GHz<=f<=40GHz |  | -47dBm/1MHz | -11.34 (NOTE 2) | 17.2 |  |
|  | 40GHz<=f<=80GHz |  | -47dBm/1MHz | -25.1 (NOTE 3) | 33.1 |  |
|  | 80GHz<f<=87GHz |  | -47dBm/1MHz | -27.1 (NOTE 2) | 33.1 |  |
| NOTE 1: Estimated SNR is calculated based on agreed influence of noise. NOTE 2: Estimated SNR is calculated based on agreed relaxation value: Estimated SNR = 6dB - relaxation. NOTE 3: Estimated SNR is calculated based on agreed relaxation value: Estimated SNR = 8dB - relaxation. |  |  |  |  |  |  |

Table B.2.2.27-1a: Uncertainty value for influence of noise for PC3, MIMO for IFF

| Test case | Frequency range | Noise floor | Minimum requirement | Estimated SNRtotal [dB/400MHz] | Relaxation | Influence of noise |
| --- | --- | --- | --- | --- | --- | --- |
| MOP-EIRP | FR2a | N/A | 22.4 dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b | N/A | 20.6 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2c | N/A | 18.7 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MPR | FR2a | -7.6dBm/400MHz | 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 15.17 (NOTE 1) | 0 | 0.13 |
|  | FR2b | -5.5dBm/400MHz | 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 11.30 (NOTE 1) | 0 | 0.31 |
|  | FR2c | -4.5dBm/400MHz | 4.2dBm/ChBW (EIRP-MPB-MPR-T(MPR)=18.7-0.5-9-5) | 8.70 (NOTE 1) | 0 | 0.55 |
| MOP-TRP | FR2a | N/A | 23 dBm/ChBW | 16.33 (NOTE 1) | 0 | 0.1 |
|  | FR2b, FR2c | N/A | 23 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
| MOP-Spherical | FR2a | N/A | 11.5 dBm/ChBW | 11.45 (NOTE 1) | 0 | 0.3 |
|  | FR2b | N/A | 8 dBm/ChBW | 6.37 (NOTE 1) | 0 | 0.9 |
| Minimum output power | FR2a | -10.6dBm/400MHz | -10 dBm | -0.46 (NOTE 1) | 5.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz | -10 dBm | -4.64 (NOTE 1) | 10.5 | 1.0 (with relaxation) |
| OFF power – TRP | FR2a | -10.6dBm/400MHz | -35dBm/ChBW | -24.54 (NOTE 2) | 30.4 | 1.0 (with relaxation) |
|  | FR2b | -5.5dBm/400MHz |  | -29.64 (NOTE 2) | 35.5 | 1.0 (with relaxation) |
| SEM | FR2a | N/A | -13dBm/1MHz | 8.14 (NOTE 1) | 0 | 0.62 |
|  | FR2b, FR2c | N/A |  | 5.86 (NOTE 1) | 0 | 1.0 |
| ACLR (CP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB 16.65dBm/ChBW (EIRP-MPB-MPR-T(MPR) =22.4-0.75-3-2) Actual lowest: 7.65dBm/ChBW (EIRP-MPB-MPR-T(MPR)=22.4-0.75-9-5) | 22.86 (with 3dB MPR) (NOTE 1) | 0 | N/A |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 16.35dBm/ChBW (EIRP-MPB-MPR-T(MPR) =20.6-0.75-2-1.5) Actual lowest: 5.85dBm/ChBW (EIRP-MPB-MPR-T(MPR)=20.6-0.75-9-5) | 21.86 (with 2dB MPR) (NOTE 1) | 0 | N/A |
| ACLR (ACP) | FR2a | -7.6dBm/400MHz | Highest testable MPR for 400MHz: 3dB -0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-3-2-17) Actual lowest: -9.35 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=22.4-0.75-9-5-17) | 5.86 (NOTE 1) | 0 | 1.0 |
|  | FR2b | -5.5dBm/400MHz | Highest testable MPR for 400MHz: 2dB 0.35dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-2-1.5-16) Actual lowest: -10.15 dBm/ChBW (EIRP-MPB-MPR-T(MPR)-ACLR=20.6-0.75-9-5-16) | 5.86 (with 2dB MPR) (NOTE 1) | 0 | 1.0 |
| General Tx spurious | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| Tx spurious Co-existence | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| Additional spurious emission | Same as SISO in Table B.2.2.27-1 |  |  |  |  |  |
| NOTE 1: Estimated SNR is calculated based on agreed influence of noise. |  |  |  |  |  |  |

Table B.2.2.27-2: Uncertainty value for influence of noise for PC1, SISO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.2 |
| MOP-TRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.27 |
| MOP-Spherical | FR2a | 0 | 0.2 |
|  | FR2b | 0 | 0.35 |
| MPR | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.3 |
| Minimum output power | FR2a | 0 | 0.3 |
|  | FR2b | 0 | 0.6 |
| OFF power – TRP | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | FR2b | Same as defined for PC3 in Table B.2.2.27-1 |  |
| SEM | FR2a | 0 | 1.81 (NOTE 1, 2) |
|  | FR2b | 0 | TBD |
| ACLR (ACP) | FR2a | 0 | 0.95 |
|  | FR2b | 0 | 0.95 |
| General Tx spurious | 6GHz <=f<12.75GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 12.75GHz<=f<66GHz | 0 | 1.08 (NOTE 1, 3) |
|  | 66GHz<=f<=80GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Tx spurious Co-existence | n260 (Aggressor band : n257, n261) | 5 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | n257, n261 (Aggressor band : n260) | 3.3 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | 23.6 GHz ≤ f ≤ 24.0GHz | 0.3 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | 57 GHz ≤ f ≤ 66GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Additional spurious emission | NS_202 (7.25GHz <=f <=12.75GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | NS_202 (12.75GHz <=f <=23.45GHz) | 13 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | NS_202, NS_203 (23.6 GHz ≤ f ≤ 24.0GHz) | 0.3 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | NS_202 (23.45GHz <=f <=40.8GHz) | 13 (Same as PC3) | 2.34 (NOTE 1, 4) |
|  | NS_202 (40.8GHz <=f <=66GHz) | 13 (Same as PC3) | 2.34 (NOTE 1, 4) |
| Rx spurious | 6GHz <=f<=20GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 20GHz<=f<=40GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 40GHz<=f<=80GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| NOTE 1: values assuming up to 6% of grid points with EIRP > 43dBm. NOTE 2: values assuming SNR = -7.9dB for points with EIRP > 43dBm / SNR = 8.14dB otherwise. NOTE 3: values assuming SNR = -5dB for points with EIRP > 43dBm / SNR = 10dB otherwise. NOTE 4: values assuming SNR = -9dB for points with EIRP > 43dBm / SNR = 6dB otherwise and considering the proposed relaxation. |  |  |  |

Table B.2.2.27-2a: Uncertainty value for influence of noise for PC1, MIMO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.2 |
| MOP-TRP | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.27 |
| MOP-Spherical | FR2a | 0 | 0.2 |
|  | FR2b | 0 | 0.35 |
| MPR | FR2a | 0 | 0.13 |
|  | FR2b | 0 | 0.3 |
| Minimum output power | FR2a | 0 | 0.15 |
|  | FR2b | 0 | 0.3 |
| SEM | FR2a | 0 | 1.81 (NOTE 1, 2) |
|  | FR2b | 0 | TBD |
| ACLR (ACP) | FR2a | 0 | 0.95 |
|  | FR2b | 0 | 0.95 |
| General Tx spurious | Same as SISO in Table B.2.2.27-2 |  |  |
| Tx spurious Co-existence | Same as SISO in Table B.2.2.27-2 |  |  |
| Additional spurious emission | Same as SISO in Table B.2.2.27-2 |  |  |

Table B.2.2.27-3: Uncertainty value for influence of noise for PC5 and PC6, SISO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | 0 | 0.13 |
| MOP-TRP | FR2a | 0 | 0.13 |
| MOP-Spherical | FR2a | 0 | 0.2 |
| MPR | FR2a | 0 | 0.13 |
| Minimum output power | FR2a | 1.4dB for 400MHz 0dB otherwise | 1 |
| OFF power – TRP | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| SEM | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| ACLR (ACP) | FR2a | 0 | 0.95 |
| General Tx spurious | 6GHz <=f<12.75GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 12.75GHz<=f<66GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 66GHz<=f<=80GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Tx spurious Co-existence | n260 (Aggressor band : n257, n258) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 23.6 GHz ≤ f ≤ 24.0GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 57 GHz ≤ f ≤ 66GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Additional spurious emission | NS_202 (7.25GHz <=f <=12.75GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | NS_202 (12.75GHz <=f <=23.45GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | NS_202, NS_203 (23.6 GHz ≤ f ≤ 24.0GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | NS_202 (23.45GHz <=f <=40.8GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | NS_202 (40.8GHz <=f <=66GHz) | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Rx spurious | 6GHz <=f<12.75GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 12.75GHz<=f<66GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |
|  | 66GHz<=f<=80GHz | Same as defined for PC3 in Table B.2.2.27-1 |  |

Table B.2.2.27-3a: Uncertainty value for influence of noise for PC5 and PC6, MIMO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | 0 | 0.13 |
| MOP-TRP | FR2a | 0 | 0.13 |
| MOP-Spherical | FR2a | 0 | 0.2 |
| MPR | FR2a | 0 | 0.13 |
| Minimum output power | FR2a | 1.4dB for 400MHz 0dB otherwise | 1 |
| ACLR (ACP) | FR2a | 0 | 0.95 |

Table B.2.2.27-4: Void

Table B.2.2.27-5: Uncertainty value for influence of noise for PC7, SISO for IFF

| Test case | Frequency range | Relaxation [dB] | Influence of noise [dB] |
| --- | --- | --- | --- |
| MOP-EIRP | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| MOP-TRP | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| MOP-Spherical | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| MPR | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Minimum output power | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |
| Carrier Leakage | FR2a | Same as defined for PC3 in Table B.2.2.27-1 |  |

### B.2.2.28	Systematic error related to beam peak search

See B.2.1.28.

The uncertainty value of systematic error related to beam peak search is estimated as below table and used across clause B.

Table B.2.2.28-1: Uncertainty value for systematic error related to beam peak search for IFF

| Power class | Uncertainty value |
| --- | --- |
| PC1, PC5, PC6 | 0.7 |
| PC3, PC7 | 0.5 |

### B.2.2.29	Influence of spherical coverage grid

See B.2.1.29.

The uncertainty value of influence of spherical coverage grid is estimated as below table and used across clause B.

Table B.2.2.29-1: Uncertainty value for influence of spherical coverage grid for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC5 | 0.13 | Actual | 1.00 | 0.13 |
| PC3, PC7 | 0.12 | Actual | 1.00 | 0.12 |
| PC6 | FFS | Actual | FFS | FFS |

### B.2.2.30	Systematic error related to EIS spherical coverage

See B.2.1.30.

The uncertainty value of systematic error related to EIS spherical coverage is estimated as below table and used across clause B.

Table B.2.2.30-1: Uncertainty value for systematic error related to EIS spherical coverage for IFF

| Power class | Uncertainty value |
| --- | --- |
| PC1, PC5 | DL power step size, 0.2 |
| PC3, PC7 | DL power step size, 0.2 |
| PC6 | DL power step size, 0.2 |

### B.2.2.31	Misalignment of DUT due to change of DUT orientation

See B.2.1.31.

The uncertainty value of misalignment of DUT due to change of DUT orientation is estimated as below table and used across clause B.

Table B.2.2.31-1: Uncertainty value for misalignment of DUT due to change of DUT orientation for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1, PC5 | 0.10 | Actual | 1.00 | 0.10 |
| PC3, PC7 | 0.10 | Actual | 1.00 | 0.10 |
| PC6 | 0.10 | Actual | 1.00 | 0.10 |

### B.2.2.32	Additional Impact of Interferer ACLR

See B.2.1.32.

The uncertainty value of additional Impact of Interferer ACLR is estimated as below table and used across clause B.

Table B.2.2.32-1: Uncertainty value for additional Impact of Interferer ACLR for IFF

| Power class | Uncertainty value |
| --- | --- |
| PC1, PC3, PC5, PC6, PC7 | 0.7 |

### B.2.2.33	Modulated Interferer uncertainty

See B.2.1.33.

The uncertainty value of modulated Interferer uncertainty is estimated as below table and used across clause B.

Table B.2.2.33-1: Uncertainty value for modulated Interferer uncertainty for IFF

| Power class | Uncertainty value | Frequency range | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC5 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |
| PC3, PC7 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |
|  | 3.6 | FR2c | Normal | 2.00 | 1.8 |
| PC6 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |

### B.2.2.34	Void

### B.2.2.35	Influence of offset antenna for blocker signal

See B.2.1.35.

### B.2.2.36	Uncertainty of the RF relative power measurement equipment

See B.2.1.36.

The uncertainty value of uncertainty of the RF relative power measurement equipment is estimated as below table and used across clause B.

Table B.2.2.36-1: Uncertainty value for uncertainty of the RF relative power measurement equipment for IFF

| Power class | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- |
| PC1 | TBD | Normal | 2 | TBD |
| PC3, PC7 | [0.4] | Normal | 2 | [0.2] |
| PC6 | TBD | Normal | 2 | TBD |

## <<< Skip unchanged sections >>>

# B.3	UE maximum output power

Following tables summarize the MU threshold for EIRP and TRP measurements for UE maximum output power. The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.3-1: MU threshold for EIRP measurement for UE maximum output power

| Power Class | Diversity | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE3) | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.08 | 5.35 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | 5.28 | 5.55 |
|  |  | 40.8GHz < f <= 44.3GHz |  |  | 6.64 | 6.78 |
| PC1 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.33 | 5.60 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | 5.40 | 5.67 |
| PC5 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.33 | 5.60 |
| PC6 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.31 | 5.58 |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.3.2-2 for PC3 and PC6 UEs (NTC), in Table B.3.2-8 for PC3 and PC6 UEs (ETC) and B.3.2-6 (NTC) and B.3.2-9 (ETC) for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. NOTE 3: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz), SISO and MBW <=100MHz. |  |  |  |  |  |  |

Table B.3-2: MU threshold for TRP measurement for UE maximum output power

| Power Class | Diversity | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE 1) | Threshold MU value for ETC [dB] (NOTE 1) |
| --- | --- | --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE3) | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.61 | 4.87 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | 4.81 | 5.07 |
|  |  | 40.8GHz < f <= 44.3GHz |  |  | 6.16 | 6.30 |
| PC1 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.64 | 4.90 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | 4.78 | 5.04 |
| PC5 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.64 | 4.90 |
| PC6 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.64 | 4.90 |
| NOTE 1: Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.3.2-2 for PC3 and PC6 UEs (NTC), in Table B.3.2-8 for PC3 and PC6 UEs (ETC) and B.3.2-6 (NTC) and B.3.2-9 (ETC) for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. NOTE 3: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz), SISO and MBW <=100MHz. |  |  |  |  |  |  |

Table B.3-3: MU threshold for Spherical coverage measurement for UE maximum output power

| Power Class | Diversity | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE2) | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.78 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | 5.38 |
|  |  | 40.8GHz < f <= 44.3GHz |  |  | 6.84 |
| PC1 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.69 |
|  |  | 32.125GHz < f <= 40.8GHz |  |  | 4.84 |
| PC5 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.69 |
| PC6 | SISO, MIMO | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | TBD |
| NOTE 1: Total Spherical coverage Expanded MU for IFF for Quiet Zone size ≤ 30cm in Tables B.3.2-4 for PC3 UEs and B.3.2-7 for PC1 and PC5 UEs. NOTE 2: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz), SISO and MBW <=100MHz. |  |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.4	UE maximum output power for modulation / channel bandwidth

Following tables summarize the MU threshold for EIRP measurements for UE maximum output power for modulation / channel bandwidth (a.k.a Maximum Power Reduction/MPR). The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.4-1: MU threshold for EIRP measurement for UE maximum output power for modulation / channel bandwidth

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE4) | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MBR - MPR – T(MPR) | 5.11 | 5.38 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.29 | 5.56 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.89 | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MPR – T(MPR) | 5.33 | 5.60 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.50 | 5.77 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MPR – T(MPR) | 5.33 | 5.60 |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.4.2-2 and Table B.4.2-3 for PC3 UEs and Table B.4.2-4 and Table B.4.2-5 for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. NOTE 3: The MU values are valid for SISO and MIMO. NOTE 4: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz), SISO and MBW <=100MHz. |  |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.7	Minimum Output power

Following tables summarize the MU threshold for EIRP measurements for Minimum Output Power. The origin MU values for different test setups can be found in following clauses.

Table B.7-1: MU threshold for EIRP measurement for Minimum output power

| Power Class | Diversity | Frequency | MBW | Power | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value For ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- | --- |
| PC1 | SISO | 23.45GHz ≤ f ≤ 32.125GHz | BW ≤ 400MHz | P = Minimum Output Power | 5.66 | 5.92 |
|  |  | 32.125GHz < f ≤ 40.8GHz |  |  | 5.96 | 6.22 |
|  | MIMO | 23.45GHz ≤ f ≤ 32.125GHz |  |  | 5.51 | 5.77 |
|  |  | 32.125GHz < f ≤ 40.8GHz |  |  | 5.66 | 5.92 |
| PC3, PC7 (NOTE2) | SISO, MIMO | 23.45GHz ≤ f ≤ 32.125GHz | BW ≤ 400MHz | P = Minimum Output Power | 6.15 | 6.41 |
|  |  | 32.125GHz < f ≤ 40.8GHz |  |  | 6.15 | 6.41 |
|  |  | 40.8GHz < f ≤ 44.3GHz |  |  | 7.34 | TBD |
| PC5 | SISO, MIMO | 23.45GHz ≤ f ≤ 32.125GHz | BW ≤ 400MHz | P = Minimum Output Power | 6.36 | 6.62 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.7.2-2 and Table B.7.2-5 for PC3 UEs and Table B.7.2-3 and Table B.7.2-6 for PC1 and PC5 UEs. NOTE 2: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz), SISO and MBW <=100MHz. |  |  |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.11	Carrier leakage

Editor’s Note: MU value analysis for PC1, 2 and 4 are not complete.

Following tables summarize the MU threshold for EIRP measurements for carrier leakage. The origin MU values for different test setups can be found in following subclauses.

Table B.11-1: MU threshold for EIRP measurement for carrier leakage

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE2) | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = 0 + MU to 0 + (MU + Uplink power control window size) dBm | 5.44 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.57 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | FFS | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.11.2-2 for PC3 UEs and FFS for PC1 UEs. NOTE 2: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz), SISO and MBW <=100MHz. |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.19	Reference Sensitivity

Following tables summarize the MU threshold for EIS measurements for Reference Sensitivity. The origin MU values for different test setups with varies parameters can be found in following subclauses.

Table B.19-1: MU threshold for EIS for Reference Sensitivity

| Power Class | Frequency | MBW | Power | Threshold MU value for NTC (NOTE 1) | Threshold MU value for ETC (NOTE 1) |
| --- | --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE2) | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.36 | 5.61 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.36 | 5.61 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.34 | 6.48 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.58 | 5.83 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.58 | 5.83 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.58 | 5.83 |
| PC6 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.56 | 5.81 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.19.2-2 for PC3 and PC6 UEs (NTC), in Table B.19.2-4 for PC3 and PC6 UEs (ETC), and Table B.19.2-5 for PC1 and PC5 UEs. NOTE 3: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz) and MBW <=100MHz. |  |  |  |  |  |

Table B.19-2: MU threshold for Spherical coverage measurement for Reference Sensitivity

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3, PC7 (NOTE2) | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.07 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.07 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.04 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.07 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.07 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.07 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.19.2-2 for PC3 UEs and Table B.19.2-3 for PC1 and PC5 UEs. NOTE 2: MU thersholds for PC7 limited to FR2a (23.45GHz <= f <= 32.125GHz) and MBW <=100MHz. |  |  |  |  |

## <<< END OF CHANGES >>>