3GPP TSG RAN WG5 Meeting #84	R5-196854r2

Ljubljana, Slovenia, 26th – 30th August 2019

Source:		Rohde & Schwarz

Title:	Discussion on Quiet Zone sizes and measurement uncertainties

Agenda Item:		5.3.5.17

Document for:	Approval

# Introduction

In the current discussions concerning FR2 measurement uncertainties and requirements two quiet zone sizes of 15 and 30 cm each are being discussed. This leads to extended discussions and an unnecessarily complex specification. In this paper we propose a way to simplify this approach.

# Discussion

## Measurement uncertainty and quiet zone size

During the FR2 testability study item in RAN4 assumptions needed to be made on the size of the quiet zone of the test system, to be able to estimate some initial measurement uncertainties and to have a common understanding of what the test system could actually look like. In these discussions it was determined that a quiet zone size of 15 cm should be used as a basis for the analysis, since this size would likely allow to measure smartphones. Later on during RAN5 discussions, several companies announced the need to test larger devices would also need to be tested, so a second quiet zone size of 30 cm was established.

As a part of the RAN5 work measurement uncertainties for FR2 systems, companies then provided measurement uncertainty contributions for test systems with either 15 or 30 cm quiet zone sizes. This lead to the introduction of two measurement uncertainty values for each test case, one for 15 cm and one for 30 cm quiet zone size.

When looking at the measurement uncertainty calculations in TR 38.903 [1], it can be seen from the tables for each test case, that the only difference in the tables for 15 und 30 cm is the actual Quality of the Quiet Zone. When comparing the overall difference in uncertainty the difference comes down to 0.1 dB, as shown in the tables below, that are extracted from TR 38.903 [1].

Table 2-1 Measurement uncertainty for maximum output power (15 cm)

| Total measurement uncertainty | Value |
| --- | --- |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.89 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] | 5.09 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.42 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.62 |

Table 2-2 Measurement uncertainty for maximum output power (30 cm)

| Total measurement uncertainty | Value |
| --- | --- |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.79 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.99 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.32 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] | 4.52 |

Observation 1: Difference between measurement uncertainties for 15 and 30 cm quiet zone sizes is only 0.1 dB.

In the current iteration of [1] for each test case two tables are maintained one for each quiet zone size. This creates an overly long specification and makes it also difficult to maintain, since changes will need to be made to multiple tables in case parameters need to be added or adapted.

Observation 2: For each TC two tables need to be maintained in the specification, creating a lot of maintenance effort.

Therefore, we propose to simplify the specification and to only use one measurement uncertainty for each TC, independent of the quiet zone size. In our understanding, the measurement uncertainty should be specified in a way that one uncertainty applies to all quiet zone sizes less or equal to 30 cm.

From a specification point of view, this would have several benefits, without having drawbacks. The effort of maintaining the specification will be reduced and the readability of the specification will also be greatly increased.

Proposal 1: Use only one measurement uncertainty in the specification for all quiet zone sizes less or equal to 30 cm based on the values for 15 cm quiet zone size (worst case).

In the Annex A an example is shown for the needed changes in TR 38.903 [1] and TS 38.521-2 [2] for the maximum output power case. The CRs associated to this discussion papers are Refs. [3] and [4].

In the measurement uncertainty definition RAN5 has so far focused only on defining uncertainties for IFF systems, so the approach should currently be limited to these systems. In case uncertainties for other system types are defined as well later on, this approach could be extended to other system types.

Proposal 2: Apply the approach only to IFF systems, since no other system type has been investigated so far. Maintain separate tables per quiet zone size for the other permitted methods.

# Conclusion

In this contribution we propose a way to simplify the measurement uncertainty specification in TR 38.903 and TS 38.521-2 by using only one measurement uncertainty for all quiet zone sizes less or equal to 30 cm for IFF systems.

Observation 1: Difference between measurement uncertainties for 15 and 30 cm quiet zone sizes is only 0.1 dB.

Observation 2: For each TC two tables need to be maintained in the specification, creating a lot of maintenance effort.

Proposal 1: Use only one measurement uncertainty in the specification for all quiet zone sizes less or equal to 30 cm based on the values for 15 cm quiet zone size (worst case).

Proposal 2: Apply the approach only to IFF systems, since no other system type has been investigated so far. Maintain separate tables per quiet zone size for the other permitted methods.

# References

- TR 38.903, V16.0.0, 3GPP TSG RAN WG5, June 2019

- TS 38.521-2, V16.0.0, 3GPP TSG RAN WG5, June 2019

- R5-196840, Update of FR2 MUs in TS 38.521-2, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #84, August 2019

- R5-196841, Update of FR2 MUs in TR 38.903, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #84, August 2019

# Annex A:

## A.1 Changes to TR 38.903:

B.3.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.2-1.

Table B.3.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID |  | Description of uncertainty contribution |  | Details in annex |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.2.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.2.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.2.3 |  |
| 4 |  | Mismatch |  | B.2.2.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.2.5 |  |
| 6 |  | Uncertainty of the RF power measurement equipment |  | B.2.2.6 |  |
| 7 |  | Phase curvature |  | B.2.2.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.2.8 |  |
| 9 |  | Random uncertainty |  | B.2.2.9 |  |
| 10 |  | Influence of the XPD |  | B.2.2.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.2.12 |  |
| 13 |  | Influence of TRP measurement grid |  | B.2.2.22 |  |
| 14 |  | Influence of beam peak search grid |  | B.2.2.23 |  |
| 15 |  | Multiple measurement antenna uncertainty |  | B.2.2.25 |  |
| 16 |  | DUT repositioning |  | B.2.2.26 |  |
| 17 |  | Influence of spherical coverage grid |  | B.2.2.29 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 |  | Mismatch |  | B.2.2.4 |  |
| 19 |  | Amplifier Uncertainties |  | B.2.2.8 |  |
| 20 |  | Misalignment of positioning System |  | B.2.2.13 |  |
| 21 |  | Uncertainty of the Network Analyzer |  | B.2.2.14 |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.2.15 |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.2.16 |  |
| 24 |  | Phase centre offset of calibration antenna |  | B.2.2.18 |  |
| 25 |  | Quality of quiet zone for calibration process |  | B.2.2.19 |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.2.20 |  |
| 27 |  | Influence of the calibration antenna feed cable |  | B.2.2.21 |  |
| 28 |  | Insertion Loss Variation |  | B.2.1.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 29 |  | Systematic error due to TRP calculation/quadrature |  | B.2.2.24 |  |
| 30 |  | Influence of noise |  | B.2.1.27 |  |
| 31 |  | Systematic error related to beam peak search |  | B.2.2.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size <= [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}, [P = maximum output power].

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.2-2 and Table B.3.2-3.

-	The uncertainty assessment for Spherical coverage is provided in Table B.3.2-4 and Table B.3.2-5.

Table B.3.2-2: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.89 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.09 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.42 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.62 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-3: Void

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1, NOTE 9) | 0.5 | Actual | 1.00 | 0.5 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 10) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1, NOTE 9) | 0.2 | Actual | 1.00 | 0.2 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |
| 30 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |
| 31 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.79 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.99 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.32 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.52 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void NOTE 8: Void NOTE 9: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size of 30 cm. NOTE 10: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

## A.2 Changes to TR 38.521-2:

Annex F:

Table F.1.2-1: Maximum Test System Uncertainty (MTSU) for transmitter tests

| Sub clause | Maximum Test System Uncertainty | Derivation of MTSU |
| --- | --- | --- |
| 6.2.1.1 UE maximum output power (EIRP) | PC3 Minimum peak EIRP, Max EIRP Quiet Zone ≤ 30 cm ±4.89 dB (FR2a) ±5.09 dB (FR2b) | MTSU = 1.00 x MU (from Table B.3-2-2, in TR 38.903) |
| 6.2.1.1 UE maximum output power (TRP) | PC3 Max TRP Quiet Zone ≤ 30 cm ±4.42 dB (FR2a) ±4.62 dB (FR2b) | MTSU = 1.00 x MU (from Table B.3-2-2, in TR 38.903) |

Table F.3.2-1: Derivation of Test Requirements (Transmitter tests)

| Sub clause | Test Tolerance (TT) | Formula for test requirement |
| --- | --- | --- |
| 6.2.1.1 UE maximum output power (EIRP) | PC3 Minimum peak EIRP IFF (Quiet Zone ≤ 30 cm) 3.18 dB (FR2a) 3.31 dB (FR2b) Max EIRP 0 dB | Minimum peak EIRP TT = 0.65 x MTSUIFF |
| 6.2.1.1 UE maximum output power (TRP) | PC3 Max TRP IFF (Quiet Zone ≤ 30 cm) 2.87 dB (FR2a) 3.00 dB (FR2b) IFF (DUT ≤ 30 cm) 2.81 dB (FR2a) 2.94 dB (FR2b) | Max TRP TT = 0.65 x MTSUIFF |

Maximum Output power TC:

6.2.1.1.5	Test requirement

The EIRP derived in step 4 and TRP derived in step 5 shall not exceed the values specified in Table 6.2.1.1.5-1 to Table 6.2.1.1.5-4.

Table 6.2.1.1.5-1: UE maximum output test requirements for power class 1

| Operating band | Max TRP (dBm) | Max EIRP (dBm) | Min peak EIRP (dBm) |
| --- | --- | --- | --- |
| n257 | 35+TT | 55 | 40.0-TT |
| n258 | 35+TT | 55 | 40.0-TT |
| n260 | 35+TT | 55 | 38.0-TT |
| n261 | 35+TT | 55 | 40.0-TT |

Table 6.2.1.1.5-2: UE maximum output test requirements for power class 2

| Operating band | Max TRP (dBm) | Max EIRP (dBm) | Min peak EIRP (dBm) |
| --- | --- | --- | --- |
| n257 | 23+TT | 43 | 29-TT |
| n258 | 23+TT | 43 | 29-TT |
| n260 |  |  |  |
| n261 | 23+TT | 43 | 29-TT |

Table 6.2.1.1.5-3: UE maximum output test requirements for power class 3 for single band UE or multiband UE declaring MBp = 0 in all FR2 bands

| Operating band | Max TRP (dBm) | Max EIRP (dBm) | Min peak EIRP (dBm) |
| --- | --- | --- | --- |
| n257 | 23+TT | 43 | 22.4-TT |
| n258 | 23+TT | 43 | 22.4-TT |
| n260 | 23+TT | 43 | 20.6-TT |
| n261 | 23+TT | 43 | 22.4-TT |

Table 6.2.1.1.5-3a: UE maximum output test requirements for power class 3 for multi band UE declaring MBp>0 in any FR2 band

| ID | Supported FR2 bands set | Test requirement (dB) (Note 1) |  |  |  | Maximum sum of MBp, ∑MBP (dB) (Note 3) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | n257 | n258 | n260 | n261 |  |  |
| 1 | n257, n258 | 22.4-TT-MBp | 22.4-TT-MBp |  |  | 1.3 |  |
| 2 | n257, n260 | 22.4-TT-MBp |  | 20.6-TT-MBp |  | 1.0 |  |
| 3 | n258, n260 |  | 22.4-TT-MBp | 20.6-TT-MBp |  | 1.0 |  |
| 4 | n258, n261 |  | 22.4-TT-MBp |  | 22.4-TT-MBp | 1.0 |  |
| 5 | n260, n261 |  |  |  |  | 0.0 | No relaxation factor allowed |
| 6 | n257, n258, n260 | 22.4-TT-MBp | 22.4-TT-MBp | 20.6-TT-MBp |  | 1.7 |  |
| 7 | n257, n258, n261 | 22.4-TT-MBp | 22.4-TT-MBp |  | 22.4-TT-MBp | 1.7 |  |
| 8 | n257, n260, n261 | 22.4-TT-MBp |  | 20.6-TT-MBp | 22.4-TT-MBp | 0.5 |  |
| 9 | n258, n260, n261 |  | 22.4-TT-MBp | 20.6-TT-MBp | 22.4-TT-MBp | 1.5 |  |
| 10 | n257, n258, n260, n261 | 22.4-TT-MBp | 22.4-TT-MBp | 20.6-TT-MBp | 22.4-TT-MBp | 1.7 |  |
| Note 1: MBp is the Multiband Relaxation factor declared by the UE for the tested band in table A.4.3.9-2 of TS38.508-2. This declaration shall fulfil the requirements in clause 6.2.1.1.3.3. Note 2: All UE supported bands needs to be tested to ensure the multiband relaxation declaration is compliant Note 3: Max allowed sum of MBp over all supported FR2 bands as defined in clause 6.2.1.1.3.3 |  |  |  |  |  |  |  |

Table 6.2.1.1.5-3b: Test Tolerance (Max TRP for Power class 3)

| Test Metric | FR2a | FR2b |
| --- | --- | --- |
| Quiet Zone ≤ 30 cm | 2.87 dB | 3.00 dB |

Table 6.2.1.1.5-3c: Test Tolerance (Min peak EIRP for Power class 3)

| Test Metric | FR2a | FR2b |
| --- | --- | --- |
| Quiet Zone ≤ 30 cm | 3.18 dB | 3.31 dB |

Table 6.2.1.1.5-4: UE maximum output power test requirements for power class 4

| Operating band | Max TRP (dBm) | Max EIRP (dBm) | Min peak EIRP (dBm) |
| --- | --- | --- | --- |
| n257 | 23+TT | 43 | 34-TT |
| n258 | 23+TT | 43 | 34-TT |
| n260 | 23+TT | 43 | 31-TT |
| n261 | 23+TT | 43 | 34-TT |